## Slide 1: STN-LIA-MPU
An OLIA spinout development
Development/Discussion LOG
Nikos Chalikias

## Slide 2: 405 nm setup with 1000mW (electric ) power
StN-LIA-MCU with a NPN transistor to drive the 1W laser
BPW34 general purpose Photodiode
We used a 405nm Laser instead an infrared
Output is high: 2000mV (direct laser beam to Photodiode
Can see consistently a 5mV variation when exhale
Will work with similar size/power/lens Infrared Laser
Need for a 2nd Photodiode for liquid tests
It is easier to multiplex and use a 2nd laser, […if this works with the liquid project]

## Slide 3: 2000nm setup
Proposed to setup a photodiode experiment only with PCBs
  This is our goal, but many little details can block the operation for …some time
Chinna proposal: to use existing instrumentation and only PCB LIA (StN-LIA-MCU), seems much better, as a 1st step
  Then we can replace:
    TEC with PCB (2nd step)
    LDD with PCB (3d step)
    Generator with PCB (4d step)
    Add a microcontroller with LCD/buttons as experiment administrator (5th step)
    Add A PC or R-Pi as experiment administrator (6th step)
    Evaluate Preamplifier performance
    Develop and evaluate, ultra low noise multi-output PSU
This is scalable and sustainable, can start now
2000nm LED and photodiode to be delivered in a couple of days

## Slide 4: StN-LIA-MCU
V1 ongoing testing, works as expected
A 2nd PCB to be assembled
Most probably the 3d PCB will be V2
V2: Fixed PCB bugs, removed expensive connectors,
Easier to use screw terminals, local SD-Card, lower cost
2 layers PCB faster to be delivered from manufacturer, clean

## Slide 5: Plan to assess LIA performance
Started with evaluating Standard Deviation [σ] of the R component DC output of commercial LIA
  Already have results for 3 COTS LIA
Propose to expand our evaluation scope with a basic setup for a QTF characterization
  Real problem
  Experience on the setup
  No need for expensive laser or optics
  Can work with 0.5 euro 32.768 RTC crystals
  No need for Preamplifiers
  Can compare with Bari electronics

## Slide 6: KNOWN ISSUES
HW and FW is based on OLIA
Works good with internal REFERENCE signal generator but not for all Frequencies
The external REFERENCE frequency is 64X the wanted Frequency, this is confusing for the user
Works well with External REFERENCE at the order of 1KHz, but can’t follow for order of 10KHz
In our FW version we have updated the multiplier ratio to 32X: so it works for 12KHz range
Not any frequency for REFERENCE is working: It has to be an integer fraction of the sampling frequency
Sampling Frequency is 200KHz
We should not have the Integer fraction issue with external REFERENCE, as the sampling rate is (..should) be controlled by the external REFERENCE!!! ….But this not resolved yet…

## Slide 7: N2O App4Farm GAS CELL
RUN LIA 2 UNITS AT THE SAME TIME AND SAME REFERENCE
TEST IF COMMON EXTERNAL REFERENCE CREATES 2 REF OUTPUTS AT CONSTANT PHASE DIFFERENCE
BNC ARE NEEDED
IN 2 WEEKS?? YES IT IS POSSIBLE
DRIVE 5V 500mA LED, IS IMPORTANT
USE EXTERNAL PSU FOR THE LED
3+ Boards needed
STATUS
  Tested Running 2 GUIs at the same time
  3D printing a plastic optics breadboard
  To mount/integrate inside a Passepartout enclosure, bottle ouitside

## Slide 8: PCBS TO BE SOLDERED

## Slide 9: GUI
For our version of GUI: Only minor updates

## Slide 10: Schematic

## Slide 11: Function Calculate() runs every 5μs
Review of code for 1st harmonic lock-in
OLIA firmware Function calculate() is called periodically every 5μs (200KHz), to
Sample input (ADC)
Calculate the Reference sinusoidal ‘virtual-sample’ (200 samples)
  The Reference frequency is: Fsample/A, where A is integer and Fsample is 200KHz
Multiply the input sample with the Reference ‘sample’ and filter
Generate the Rectangular Reference frequency for the D.U.T.
  This is in phase with the Internally used sinusoidal Reference
0 to 199
For 1KHz: 200 samples per REF period
Input is sampled 200 times for the 1ms REF period (1KHz)
Sin (0 to 2π) in 200 steps

## Slide 12: Sinus calculation

## Slide 13: Possible upgrades
Sync the sampling
Use main/aux SNT-LIA-MPU boards with updated firmware
Aux LIA board sampling interrupt is driven from the main
  Main outputs the sampling frequency on a new output pin
  Aux receives the sampling frequency in a new input pin
Aux sampling interrupt(200KHz) is generated from the input pin, not from internal timer
Sample an external sinusoidal reference
Use a 2nd ADC analogue input pin
Apply external sinusoidal external Reference to the new ADC analogue input
Update the Firmware
  Use the external ADC sampled sinusoidal Reference, instead the internally generated
Monitor the internally generated sinusoidal Reference
Add a new DAC chip and connect external oscilloscope for monitoring
  Alternatively, export the sinusoidal Reference to a PWM output, without using a new DAC
Update the firmware to export the sinusoidal Reference to DAC, or the PWM output
Main
LIA
AUX
LIA
200KHz
ADC0
ADC1
ADC0
ADC1

## Slide 14: LIA Sampling interrupt: Upgrade for 2nd Input
ADC0
ADC1
FILTER
X
INPUT SAMPLE
2nd INPUT SAMPLE
INTERNAL SINUSOIDAL
REFERENCE GENERATOR
RECTANGULAR
REFERENCE GENERATOR
DUT
2 OUTPUTS
SYNCED
Ref=sin(2π* (1->200)/200 )
SINUSOIDAL REFERENCE ‘VIRTUAL SAMPLES’
200KHz timer Interrupt (5μs)
Fsample = 200KHz (5μs)
samplesPerPeriod = 200
Reference = 1KHz (1ms)
Serial port
DC
200KHz SYNC OUT
INTERNAL SINUSOIDAL REF EXPORT
1 KHz
200KHz SYNC INT IN
BLUE=CURRENT
GREEN=UPGRADE
100 interrupts HIGH
100 interrupts LOW
1 KHz
TO OSCILLOSCOPE
INTERNAL
Upgraded GUI protocol
GUI
GRAPHICAL
USE
INTERFACE

## Slide 15: LIA Sampling interrupt: Upgrade for external Reference
ADC0
ADC1
X
INPUT SAMPLE
REFERENCE SAMPLE
INTERNAL SINUSOIDAL
REFERENCE GENERATOR
RECTANGULAR
REFERENCE GENERATOR
D.U.T.
SYNCED
Ref=sin(2π* (1->200)/200 )
SINUSOIDAL REFERENCE
‘VIRTUAL SAMPLES’
200KHz timer Interrupt (5μs)
Fsample = 200KHz (5μs)
samplesPerPeriod = 200
Reference = 1KHz (1ms)
Serial port GUI protocol
DC
200KHz SYNC OUT
INTERNAL SINUSOIDAL REF EXPORT
1 KHz
200KHz SYNC INT IN
BLUE=CURRENT
GREEN=UPGRADE
EXT
REF
100 interrupts HIGH
100 interrupts LOW
1 KHz
TO OSCILLOSCOPE
EXTERNAL
INTERNAL
To aux LIA
From main LIA
GUI
GRAPHICAL
USE
INTERFACE
FILTER

## Slide 16: Higher Level Diagram for Orthogonal LIA
Appl. Sci. 2022, 12, 6431

## Slide 17: LIA Sampling interrupt, Includes COS : Upgrade for external Reference
ADC0
ADC1
X
INPUT SAMPLE
REFERENCE SAMPLE
INTERNAL SINUSOIDAL
REFERENCE GENERATOR
RECTANGULAR
REFERENCE GENERATOR
D.U.T.
SYNCED
Ref=sin(2π* (1->200)/200 )
SINUSOIDAL REFERENCE
‘VIRTUAL SAMPLES’
200KHz timer Interrupt (5μs)
Fsample = 200KHz (5μs)
samplesPerPeriod = 200
Reference = 1KHz (1ms)
Serial port GUI protocol
DC
200KHz SYNC OUT
INTERNAL SINUSOIDAL REF EXPORT
1 KHz
200KHz SYNC INT IN
BLUE=CURRENT
GREEN=UPGRADE
EXT
REF
100 interrupts HIGH
100 interrupts LOW
1 KHz
TO OSCILLOSCOPE
EXTERNAL
INTERNAL
To aux LIA
From main LIA
GUI
GRAPHICAL
USE
INTERFACE
X
FILTER
COSINE REFERENCE
‘VIRTUAL SAMPLES’
INTERNAL COSINE
REFERENCE GENERATOR

## Slide 18: LIA Sampling interrupt, with cos and circular buffer: Upgrade for external Reference
ADC0
ADC1
X
INPUT SAMPLE
REFERENCE SAMPLE
INTERNAL SINUSOIDAL
REFERENCE GENERATOR
RECTANGULAR
REFERENCE GENERATOR
D.U.T.
SYNCED
SINUSOID REFERENCE
‘VIRTUAL SAMPLES’
200KHz μC timer Interrupt (5μs)
Fsample = 200KHz (5μs)
samplesPerPeriod = 200
Reference = 1KHz (1ms)
X,Y,R,φ,N => Serial port GUI protocol
200KHz SYNC OUT
1 KHz
200KHz Sampling SYNC Input
BLUE=CURRENT
GREEN=UPGRADE
EXT
REF
100 interrupts HIGH
100 interrupts LOW
EXTERNAL
INTERNAL
To aux LIA
Ext Generator
Or Other LIA
X
FILTER
SAMPLED
RESULTS
COSINE REFERENCE
‘VIRTUAL SAMPLES’
INTERNAL COSINE
REFERENCE GENERATOR
CIRC
BUFFER
INTERNAL
SIN
COS
POST
PROCESS
Χ,Y
Ref=sin(2π* (1->200)/200 )
Ref=Cos(2π* (1->200)/200 )
Χ
Y
VALUE 0 TO 1
VALUE 0 TO 1

## Slide 19: EXTENDED FUNCTIONALITY

## Slide 20: StN-LIA-E: Extended Functionality
Frequency
Control
Sampling Interrupt
samplesPerPeriod
200
1KHz
GENERATOR
New
Reference Frequency
1000
200 000
Sampling
Frequency
GUI
Updated
D.U.T.
STN-LIA-MPU
aux mode
I2C
Or
Human Operator
ADC

## Slide 21: StN-LIA-E: Two Synchronized Lock In Amplifiers
Timer interrupt
Samples
Per Period
200
1KHz
Reference
Frequency
1000
200 000
Sampling
Frequency
GUI
D.U.T.
2 OUTPUTS
STN-LIA-MPU main Mode
Timer
STN-LIA-MPU aux Mode
GUI
Sampling Interrupt
from external sync Input
200
1000
200 000
ADC
ADC
Sampling Interrupt
from Internal timer

## Slide 22: LIA Sampling interrupt, Control with external Precise generator
Any Frequency, including QEPAS resolution
Precise [0.0001Hz] Frequency Resolution (0.01Hz / 100)
Can be used as External Generator for a Bench Lock In Amplifier
Synchronized operation of multiple STN-LIAs [solves UCC setup]
With a small addition the PCB setup can produce sinusoid output in the audio Freq Range

## Slide 23: Synced Input ADC channels and Circular Buffer
Synchronized reading of 2 ADC channels and circular buffer for the Reference
The circular buffer for the sinusoid Reference signal Input, is used to create the Cosine of the Reference
[0 to 199] for 1KHz and Sampling rate 200 KHz

## Slide 24: Sin and Cos implementation in C++
COS (quadrature harmonic)
Cosine is 1π (200/4 for 1KHz) later in the Circular Buffer
1*π
SIN (in phase harmonic)

## Slide 25: Requirements discussion
We're looking for two signal channels driven by a single modulation source, ideally produced by the LIA, to modulate a single LED.
From the LED, we get two signals by splitting the light into two separate channels that are individually detected after the light passes through the sample. These need to be separately lockin amplified.
We would ideally want to get a sensitive measurement of the channels individually, or maybe just the ratio of the channels if that was more sensitive and could be determined directly.
I think this means:
one modulation source (in the LIA)
two input channels
lockin signal processing for two channels . It might be possible to do the two channels from the same op amp etc., but maybe not.
After 31/12/2024 call
Must measure 2 outputs from the D.U.T. in parallel
  Implementation with 2 LIA boards that use external reference
  Implementation with 2 signal inputs on the same LIA board, with internal reference
Lower priority: Calibrate by measuring the PHASE Φ, at different frequencies1KHz to 10KHz

## Slide 26: UCC Requirements discussion: 31/12/2024 call
After 31/12/2024 call
Must measure in parallel, 2 outputs from the D.U.T.
  Implementation with 2 LIA boards that use external reference
    This option does not need extensive GUI updates
  Implementation with 2 signal inputs on the same LIA board, uses only internal reference
    Need major rework for the GUI
Lower priority: Calibrate by measuring the PHASE φ, at different frequencies1KHz to 10KHz

## Slide 27: Two PCBs tested in synchronization
aux LED is ON
Both PCBs have the same signal input (Potentiometer adjusted to 16mV)
The Reference signal is created from the main PCB and drives to Potentiometer
OPT101 and LED removed to avoid ambient light interference with the signal

## Slide 28: Two PCBs tested in synchronization
Do we need R calibration Offset (13.762mV -13.52046mV here),
So R appears exactly the same for a common input?
It would appear similar to the
Phase offset BOX

## Slide 29: Two PCBs tested in synchronization
main PCB sync out
drives aux PCB
aux PCB sync out
The 2 PCBs are synchronized
The aux PCB interrupt (sampling) has 100ns delay
The phase difference between the main and aux sampling is stable (100ns)

## Slide 30: WRONG SETUP
SETUP as reported at 29 April 2025:
main (left)Board runs new firmware
aux board (right, aka upper) runs old Firmware
Running 2 instances of the OLD GUI
The visible results
“Lock Failure” message
R, Theta on right board seems inconsistent with very low values
Nikos notes:
The setup is obviously wrong!
Slide 20 (StN-LIA-E: Two Synchronized Lock In Amplifiers) thick RED ARROW is not implemented with old firmware
  The [right, upper] aux board is not synchronized, as per slides 20,21,27,28,29
aux board
main board

## Slide 31: new GUI (unconnected to boards)
SETUP as reported at 29 April 2025:
main (left)Board runs new firmware
aux board (right, aka upper) runs old Firmware
Running 2 instances of the OLD GUI
The visible results
“Lock Failure” message
R, Theta on right board seems inconsistent with very low values
Nikos notes:
The setup is obviously wrong!
Slide 20 (StN-LIA-E: Two Synchronized Lock In Amplifiers) thick RED ARROW is not implemented with old firmware
  The [right, upper] aux board is not synchronized, as per slides 20,21,27,28,29
aux board
main board

## Slide 32: OPT101 SETUP
Expand with
2nd OPT101 PCB
2nd LED ???

## Slide 33: QUESTIONS